## Slide 1: Ortho Vision Validation and Operation in RCI
Aisling Costelloe
Senior Medical Scientist
RCI Laboratory
Aisling.Costelloe@ibts.ie

## Slide 2: Ortho Vision Validation and Operation in RCI
Role of RCI laboratory
Analyser requirements
Validation Process
Ortho Vision Validation – PQ
Validation Outcomes
Validation Deviations/Findings
Interface testing
Ortho Vision- Go Live
Ortho Vision Operation

## Slide 3: Role of the RCI Laboratory
  Hospital Blood Bank (3 institutions)
  Red Cell Immunohaematology referral laboratory
        2,500 samples per year (92% are complex in some way)
      Antenatal testing
      Investigation of complex patient samples
      Red Cell Importation
      Resolution of ABO/Rh anomalies
      Research
      Publish novel, unusual and interesting findings.
Notes:     Crossmatch and issue units for complex cases – which hospitals are unable to solve.
    Undertake complex investigations involving rare red cell antigens and antibodies.

## Slide 4: Analyser requirements
Blood Grouping & Rh/K phenotyping
Perform routine antibody ID (single specificities e.g. anti-D, anti-c
Perform antibody titrations
Perform red cell phenotyping
Perform compatibility testing
Overall perform more automated work to free bench for research/project work

## Slide 5: Validation Process
URS/Tender Process
Validation Plans (IQ/OQ/PQ)
+
Writing SOPs
+
Development & Testing of LIS/Ortho Vision Interface
+
Staff training
Go live and post performance monitoring for 3 months

## Slide 6: Ortho Vision Validation
IQ
  Software installation verification
  Hardware installation verification
  System Configuration
OQ
  Standard Operating procedures verification
  Operator qualification
  Installation verification by vendor
  Operational verification
  Consumables loading and verification
  Alarm/ Antivirus verification
  Backup and recovery verification

## Slide 7: Ortho Vision Validation- PQ
Each assay tested under the following headings:
  Sensitivity (including Limit of Detection)
  Repeatability
  Test and Process Robustness (Cross contamination)
  Reproducibility
    Intra assay Reproducibility (Repeatability)
    Inter assay Reproducibility / Parallel
Notes: Repeatability = one time point (repeat 10 times in one day)
Intra assay reproducibility = 3 different time points (Day 1, Day 4, Day 7)

## Slide 8: Ortho Vision Validation- PQ
Assays validated:
ABO/Rh Typing
Antibody Screening
Rh/K Phenotyping
Antibody Identification
Compatibility Testing
Direct Antiglobulin Test (DAT)
Extended Phenotyping
(Duffy, Kidd, MNS, k, Cw, Kpa, A1, P1, Lea, Leb, Lua)
Notes: ABO/Rh Typing Extended pheno 25 (10,10 & 5)
  Tested each ABO/Rh group 36
  Tested D variant samples (e.g. Weak D, DVI)
  Tested samples at day 7
  DAT positive samples
  Neonate samples
  ABO anomaly samples (e.g. A2, A3, Ax)

## Slide 9: Validation Outcomes
Assay | Outcome
ABO/Rh Typing | PASS
Antibody Screening | PASS (with deviations*)
Rh/K Phenotyping | PASS
Antibody Identification | PASS (with deviations*)
Compatibility testing | PASS
Direct Antiglobulin Test (DAT) | FAIL
Extended Phenotyping (Duffy, Kidd, MNS, k, Cw, Kpa, A1, P1, Lea, Leb, Lua) | PASS (with deviations*)
Notes: Deviations = didn’t completely comply with acceptance criteria?

## Slide 10: Validation Deviations/findings
Assay | Deviations/findings | Action by RCI
ABO/Rh Typing | None | N/A
Antibody Screening | 1.Weak anti-E not detected 2. Enzyme only anti-E not detected | Performed LoD testing using anti-E sample The detection of weak antibodies is a known serological limitation. Weak antisera controls are in place for antibody screen testing.
Rh/K Phenotyping | None | N/A
Antibody Identification | 1. Unable to perform ENZ-IAT technique on Ortho Vision 2. Weak anti-E not detected 3. Cross-contamination occurred on three separate occasions | 1. Perform ENZ technique on Ortho Vision as opposed to ENZ-IAT. 2. As above 3. a. Daily probe maintenance blocked b. Fresh vial of BSA used daily c. Vials of reagent cells are numbered to avoid potential carryover
Notes: Anti-E (different samples from same patient). Xcontamination – with anti-CD38, involving antisera controls and anti-Jka antisera when titred
Limit of Detection (Antibody Screen):
Anti-E (Titre: 64) – OV LoD=256
Anti-E (Titre:16) – OV LoD=64
Limit of Detection (Antibody Identification):
Anti-E (Titre:256) – OV LoD= 1024
Anti-E (Titre:1) – OV LoD= 128
ENZ-IAT – expand , background noise etc.

## Slide 11: Validation Deviations/findings
Assay | Deviations/findings | Action by RCI
Compatibility testing | Incompatibility not detected between Fya+ donor cells and patient sample containing anti-Fya. | Incompatibility not detected by manual BioRad. Incompatibility was not detected due to sample condition. No corrective action required.
Direct Antiglobulin Test (DAT) | False negative results obtained with Ortho Vision for EQA samples only | RCI did not progress with DAT testing on Ortho Vision
Extended Phenotyping (Duffy, Kidd, MNS, k, Cw, Kpa, A1, P1, Lea, Leb, Lua) | Anti-N negative control giving positive results | RCI did not progress with N phenotype
Notes: DAT false negative
EQA
Donor sample via ADG

## Slide 12: Interface testing
LIS = eTraceline (MAK)
Development of result transfer between IBTS-MAK-Ortho/Brennans
Multi-departmental

## Slide 13: Ortho Vision- Go Live
Go live date: March 2021
4 x test samples ran initially which covered all testing profiles
Issues on go live:
  Interface related only
Training
  Basic training provided for all staff in advance
  Priority given to on-call staff

## Slide 14: Ortho Vision Operation
Advantages
  Automated assays for previously manual methods
    Titration
    ID
    Use of NHSBT panel for exclusions
    Phenotyping
  Bidirectional interface
    Transfer of reaction grades, reagent lot #/ expiry date direct to patient file on LIS
  QC tracking
  Product & Lot No. easily tracked
  Usage data available

## Slide 15: Ortho Vision Operation
Disadvantages
  TAT
    Multiple different test profiles run at once
  Sample Volume (system uses 50ul of plasma for Ab ID)
  ENZ-IAT not available
  Unable to modify sample type
  Unable to register anti-Fya as a control (Antibody screen/ID gives discrepant results and thus fails QC)

## Slide 16: Ortho Vision Operation
Incidents since go live:
Errors with interface
Weak Anti-D not detected by Ortho Vision – detected by manual BioRad technique
  All negative results by ENZ and IAT are confirmed with manual ENZ-IAT
Phenotyping controls
  M Type [Error: “?” or “TFC”]
  Fyb Type [<2+ reactions i.e. not acceptable]

## Slide 17
THANK YOU!